| ISSN: 2231-3656 Print: 2231-3648 | International Journal of Pharmacy and Industrial Research (IJPIR) IJPIR ¦Vol.13 ¦ Issue 4 ¦ Oct - Dec -2023 www.ijpir.com DOI : https://doi.org/10.61096/ijpir.v13.iss4.2023.187-195 |
| --- | --- |

Research

Simultaneous Estimation Of New Analytical Method Development And Validation Of Dabrafenib And Trametinib By High Performance Liquid Chromatography

Bramhadevu Swati Chandra Shekar*, R. Mounika, K.Chaitanya Prasad, B.Sudhakar, N.Sravya, R.Madhulika

1Department of Pharmaceutical Analysis, Samskruti College Of Pharmacy In Ghatkesar, Telangana. 501301.

*Author for Correspondence: Bramhadevu Swati Chandra Shekar

Email: chandrasekhar2729@gmail.com

|  | Abstract |
| --- | --- |
| Published on: 20 Oct 2023 | An accurate, precise, simple, efficient and reproducible, isocratic Reversed Phase-High Performance Liquid Chromatography (RP-HPLC) method was developed and validated for the simultaneous estimation of Dabrafenib and Trametinib in bulk and combined pharmaceutical tablet dosage forms. Dabrafenib and Trametinib were separated by using a Symmetry ODS C18 (4.6mm×150mm) 5µm Particle Size; Waters Alliance e2695 HPLC system with 2998 PDA detector and the mobile phase contained a mixture of Methanol: 0.1% Orthophosphoric acid (64:36% v/v). The flow rate was set to 1ml/min with the responses measured at 224nm. The retention time of Dabrafenib and Trametinib was found to be 2.808min and 3.880min respectively with resolution of 5.68. Linearity was established for Dabrafenib and Trametinib in the range of 20-100µg/ml for Dabrafenib and 60-140µg/ml for Trametinib with correlation coefficient 0.999. The percentage recovery was found to be is 100.30% for Dabrafenib and 100.21% for Trametinib respectively. Validation parameters such as specificity, linearity, precision, accuracy and robustness, limit of detection (LOD) and limit of quantitation (LOQ) were evaluated for the method according to the International Conference on Harmonization (ICH) Q2 R1 guidelines. The developed method was successfully applied for the quantification of bulk and active pharmaceutical ingredient present and in combined tablet dosage form. |
| Published by: DrSriram Publications |  |
| 2023¦ All rights reserved. Creative Commons Attribution 4.0 International License. |  |
|  | Keywords: Dabrafenib and Trametinib, RP-HPLC, Validation, Accuracy, Precision |

INTRODUCTION

Analytical chemistry1 is the branch of chemistry involved in separating, identifying and determining the relative amounts of the components making up a sample of matter. It is mainly involved in the qualitative identification or detection of compounds and the quantitative measurement of the substances present in bulk and pharmaceutical preparation.

Measurements of physical properties of analytes such as conductivity, electrode potential, light absorption or emission, mass to charge ratio, and fluorescence, began to be used for quantitative analysis of variety of inorganic and biochemical analytes. Highly efficient chromatographic and electrophoretic techniques began to replace distillation, extraction and precipitation for the separation of components of complex mixtures prior to their qualitative or quantitative determination. These newer methods for separating and determining chemical species are known collectively as instrumental methods of analysis. Most of the instrumental methods fit into one of the three following categories viz spectroscopy, electrochemistry and chromatography

Advantages of instrumental methods

- Small samples can be used

- High sensitivity is obtained

- Measurements obtained are reliable

- Determination is very fast

- Even complex samples can be handled easily

Limitations of instrumental methods

- An initial or continuous calibration is required

- Sensitivity and accuracy depends on the instrument

- Cost of equipment is large

- Concentration range is limited

- Specialized training is needed

- Sizable space is required

Principle types of chemical instrumentation:

Spectrometric techniques

- Ultraviolet and visible Spectrophotometry

- Fluorescence and phosphorescence Spectrophotometry.

- Atomic Spectrometry (emission and absorption)

- Infrared Spectrophotometry

- Raman Spectroscopy

- X-Ray Spectroscopy

- Radiochemical Techniques including activation analysis

- Nuclear Magnetic Resonance Spectroscopy

- Electron Spin Resonance Spectroscopy

Electrochemical techniques

- Potentiometry

- Voltametry

- Voltametric Techniques

- Amperometric Techniques

- Colorimetry

- Electrogravimetry

- Conductance Techniques

Chromatographic techniques

- Gas Chromatography

- High Performance Liquid Chromatography

- High Performance Thin Layer  Chromatography

# Miscellaneous techniques

- Thermal Analysis

- Mass Spectrometry

- Kinetic Techniques

- Hyphenated techniques

- GC-MS (Gas Chromatography – Mass Spectrometry)

- GC-IR (Gas Chromatography – Infrared Spectroscopy)

- MS-MS (Mass Spectrometry – Mass Spectrometry)

1.1 INTRODUCTION TO HPLC

HPLC3 is a type of liquid chromatography that employs a liquid mobile phase and a very finely divided stationary phase. In order to obtain satisfactory flow rate liquid must be pressurized to a few thousands of pounds per square inch.

The rate of distribution of drugs between Stationary and mobile phase is controlled by diffusion process. If diffusion is minimized faster and effective separation can be achieved. The techniques of high performance liquid chromatography are so called because of its improved performance when compared to classical column chromatography advances in column chromatography into high speed, efficient, accurate and highly resolved method of separation.

For the recent study Clonazepam and Propranolol was selected for estimation of amount of analyte present in formulation and bulk drug. The HPLC method is selected in the field of analytical chemistry, since this method is specific, robust, linear, precise and accurate and the limit of detection is low and also it offers the following advantages

- Speed many analysis can be accomplished in 20min (or) less.

- Greater sensitivity (various detectors can be employed).

- Improved resolution (wide variety of stationary phases).

- Re usable columns (expensive columns but can be used for many analysis).

- Ideal for the substances of low viscosity.

- Easy sample recovery, handling and maintenance.

- Instrumentation leads itself to automation and quantification (less time and less labour).

- Precise and reproducible.

- Integrator itself does calculations.

- Suitable for preparative liquid chromatography on a much larger scale.

MATERIALS AND METHODS

Dabrafenib from Sura labs, Trametinib from Sura labs, Water and Methanol for HPLC from LICHROSOLV (MERCK). Acetonitrile for HPLC from Merck, Phosphate buffer from Sura labs.

HPLC METHOD DEVELOPMENT

TRAILS

Preparation of standard solution

Accurately weigh and transfer 10 mg of Dabrafenib and Trametinib  working standard into a 10ml of clean dry volumetric flasks add about 7ml of Methanol and sonicate to dissolve and removal of air completely and make volume up to the mark with the same Methanol.

Further pipette 0.6ml of Dabrafenib and 1ml of Trametinib from the above stock solutions into a 10ml volumetric flask and dilute up to the mark with Methanol.

Procedure

Inject the samples by changing the chromatographic conditions and record the chromatograms, note the conditions of proper peak elution for performing validation parameters as per ICH guidelines.

OPTIMIZED CHROMATOGRAPHIC CONDITIONS

Instrument used	:	Waters Alliance 2695 HPLC with PDA Detector 996 model.

Temperature 	           :	38ºC

Column            	: 	Symmetry ODS C18 (4.6mm×150mm) 5µm Particle Size

Mobile phase		:	Methanol: 0.1% Orthophosphoric acid (64:36% v/v)

Flow rate		: 	1ml/min

Wavelength		:	224nm

Injection volume	: 	20µl

Run time 		: 	7.0minutes

VALIDATION

PREPARATION OF MOBILE PHASE

Preparation of mobile phase

Accurately measured 640ml of Acetonitrile (64%) of and 360ml of HPLC Water (36%) were mixed and degassed in a digital ultra sonicater for 15 minutes and then filtered through 0.45 µ filter under vacuum filtration.

Diluent Preparation

The Mobile phase was used as the diluent.

RESULTS AND DISCUSSION

Optimized Chromatogram (Standard)

Mobile phase           :   Methanol: 0.1% Orthophosphoric acid (64:36% v/v)

Column                   :   Symmetry ODS C18 (4.6mm×150mm) 5µm Particle Size

Flow rate                 :   1 ml/min

Wavelength             :   224 nm

Column temp          :   38ºC

Sample Temp          :   Ambient

Injection Volume    :   20 µl

Run time	    :  7 minutes

Fig 1: Chromatogram for Trail 5

Table 1: Peak Results for Trail 5

| S. No | Peak name | Rt | Area | Height | USP Resolution | USP Tailing | USP plate count |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Dabrafenib | 2.808 | 65258 | 4326 |  | 1.08 | 5685.4 |
| 2 | Trametinib | 3.880 | 8659854 | 659823 | 5.68 | 1.42 | 6895.7 |

From the above chromatogram it was observed that the Dabrafenib and Trametinib peaks are well separated and they shows proper retention time, resolution, peak tail and plate count. So it’s optimized trial.

Retention time of Dabrafenib–2.808min

Retention time of Trametinib – 3.880 min

Assay (Standard)

Table 2: Showing assay standard Results

| S.No. | Name | Rt | Area | Height | USP Resolution | USP Tailing | USP plate count | Injection |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Dabrafenib | 2.813 | 65684 | 4365 |  | 1.08 | 5632.4 | 1 |
| 2 | Trametinib | 3.886 | 8659824 | 659824 | 5.69 | 1.42 | 6859.2 | 1 |
| 3 | Dabrafenib | 2.813 | 65985 | 4329 |  | 1.09 | 5682.3 | 2 |
| 4 | Trametinib | 3.886 | 8645872 | 658266 | 5.68 | 1.43 | 6824.1 | 2 |
| 5 | Dabrafenib | 2.813 | 65784 | 4426 |  | 1.08 | 5692.8 | 3 |
| 6 | Trametinib | 3.886 | 8657847 | 6589412 | 5.69 | 1.43 | 6895.4 | 3 |

Assay (Sample)

Table 3: Showing assay sample results

| S.No. | Name | Rt | Area | Height | USP Resolution | USP Tailing | USP plate count | Injection |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Dabrafenib | 2.799 | 66859 | 4458 |  | 1.09 | 5785.4 | 1 |
| 2 | Trametinib | 3.863 | 8756854 | 669585 | 5.69 | 1.43 | 6956.7 | 1 |
| 3 | Dabrafenib | 2.799 | 66258 | 4462 |  | 1.10 | 5789.5 | 2 |
| 4 | Trametinib | 3.861 | 8769582 | 663598 | 5.68 | 1.44 | 6945.2 | 2 |
| 5 | Dabrafenib | 2.799 | 66435 | 4438 |  | 1.09 | 5784.1 | 3 |
| 6 | Trametinib | 3.863 | 8754985 | 668548 | 5.69 | 1.44 | 6927.7 | 3 |

Sample area        Weight of standard     Dilution of sample     Purity      Weight of tablet

%ASSAY =    ___________ ×   ________________ × _______________×  _______× ______________×100

Standard area      Dilution of standard    Weight of sample       100             Label claim

The % purity of Tinidazole and Diloxanide in pharmaceutical dosage form was found to be 99.4 %.

LINEARITY

CHROMATOGRAPHIC DATA FOR LINEARITY STUDY

Tinidazole

Fig 2: calibration graph for Dabrafenib

Linearity Results: (for Dabrafenib)

| S.No. | Linearity Level | Concentration (ppm) | Area |
| --- | --- | --- | --- |
| 1 | I | 20 | 24759 |
| 2 | II | 40 | 47859 |
| 3 | III | 60 | 70898 |
| 4 | IV | 80 | 93985 |
| 5 | V | 100 | 116698 |
| Correlation Coefficient |  |  | 0.999 |

Correlation coefficient should be not less than 0.999.

Linearity Results: (for Trametinib)

Fig 3: Calibration graph for Trametinib

| S.No. | Linearity Level | Concentration (ppm) | Area |
| --- | --- | --- | --- |
| 1 | I | 60 | 4928578 |
| 2 | II | 80 | 6687842 |
| 3 | III | 100 | 8389878 |
| 4 | IV | 120 | 10085847 |
| 5 | V | 140 | 11769854 |
| Correlation Coefficient |  |  | 0.999 |

- Correlation coefficient should be not less than 0.99.

REPEATABILITY

Table 4: Results of method precision for Dabrafenib

| S.No. | Name | Rt | Area | Height | USP plate count | USP Tailing |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Dabrafenib | 2.808 | 65898 | 4365 | 5682.2 | 1.08 |
| 2 | Dabrafenib | 2.808 | 65487 | 4375 | 5628.6 | 1.09 |
| 3 | Dabrafenib | 2.808 | 65324 | 4395 | 5649.7 | 1.08 |
| 4 | Dabrafenib | 2.808 | 65982 | 4328 | 5638.4 | 1.09 |
| 5 | Dabrafenib | 2.808 | 65248 | 4371 | 5698.3 | 1.08 |
| 6 | Dabrafenib | 2.808 | 65734 | 4391 | 5682.7 | 1.09 |
| Mean |  |  | 65612.17 |  |  |  |
| Std. Dev |  |  | 304.8425 |  |  |  |
| % RSD |  |  | 0.464613 |  |  |  |

Table 5: Results of method precision for Trametinib

| S.No. | Name | Rt | Area | Height | USP plate count | USP Tailing | USP Resolution |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Trametinib | 3.880 | 8659824 | 658784 | 6859.4 | 1.42 | 5.68 |
| 2 | Trametinib | 3.880 | 8658547 | 657489 | 6824.6 | 1.43 | 5.69 |
| 3 | Trametinib | 3.880 | 8659824 | 652368 | 6829.3 | 1.42 | 5.68 |
| 4 | Trametinib | 3.880 | 8659875 | 658745 | 6892.7 | 1.43 | 5.69 |
| 5 | Trametinib | 3.880 | 8658745 | 658213 | 6875.2 | 1.42 | 5.68 |
| 6 | Trametinib | 3.880 | 8659862 | 652354 | 6859.8 | 1.42 | 5.69 |
| Mean |  |  | 8659446 |  |  |  |  |
| Std. Dev |  |  | 623.2924 |  |  |  |  |
| % RSD |  |  | 0.007198 |  |  |  |  |

- %RSD for sample should be NMT 2.

- The %RSD for the standard solution is below 1, which is within the limits hence method is precise.

Intermediate precision

Table 6: Results of Intermediate precision for Dabrafenib:

| S.No. | Name | Rt | Area | Height | USP plate count | USP Tailing |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Dabrafenib | 2.808 | 66895 | 4468 | 5784.2 | 1.09 |
| 2 | Dabrafenib | 2.808 | 66986 | 4523 | 5835.1 | 1.09 |
| 3 | Dabrafenib | 2.808 | 66258 | 4475 | 5864.4 | 1.10 |
| 4 | Dabrafenib | 2.808 | 66457 | 4514 | 5864.6 | 1.09 |
| 5 | Dabrafenib | 2.808 | 66539 | 4489 | 5784.9 | 1.10 |
| 6 | Dabrafenib | 2.808 | 66298 | 4565 | 5748.5 | 1.10 |
| Mean |  |  | 66572.17 |  |  |  |
| Std. Dev |  |  | 304.536 |  |  |  |
| % RSD |  |  | 0.457452 |  |  |  |

Table 7: Results of Intermediate precision for Trametinib

| S.No. | Name | Rt | Area | Height | USP plate count | USP Tailing | USP Resolution |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Trametinib | 3.882 | 8758568 | 669583 | 6982.4 | 1.43 |  |
| 2 | Trametinib | 3.882 | 8756982 | 665984 | 6935.3 | 1.44 | 5.69 |
| 3 | Trametinib | 3.882 | 8746925 | 665345 | 6984.7 | 1.44 |  |
| 4 | Trametinib | 3.882 | 8723654 | 665325 | 6952.8 | 1.43 | 5.70 |
| 5 | Trametinib | 3.882 | 8754982 | 669852 | 6898.9 | 1.44 |  |
| 6 | Trametinib | 3.882 | 8754698 | 665874 | 6976.5 | 1.43 | 5.69 |
| Mean |  |  | 8749302 |  |  |  |  |
| Std. Dev |  |  | 13188.56 |  |  |  |  |
| % RSD |  |  | 0.150738 |  |  |  |  |

- %RSD of five different sample solutions should not more than 2.

- The %RSD obtained is within the limit, hence the method is rugged.

ACCURACY

Table 8: Accuracy (recovery) data for Dabrafenib

| %Concentration (at specification Level) | Area | Amount Added (mg) | Amount Found (mg) | % Recovery | Mean Recovery |
| --- | --- | --- | --- | --- | --- |
| 50% | 35921.67 | 30 | 30.134 | 100.446% | 100.30% |
| 100% | 70894.33 | 60 | 60.205 | 100.341% |  |
| 150% | 105654.7 | 90 | 90.093 | 100.103% |  |

- The % Recovery for each level should be between 98.0 to 102.0%.

Table 9: Accuracy (recovery) data for Trametinib

| % Concentration (at specification Level) | Area | Amount Added (mg) | Amount Found (mg) | % Recovery | Mean Recovery |
| --- | --- | --- | --- | --- | --- |
| 50% | 4276302 | 50 | 50.208 | 100.416% | 100.21% |
| 100% | 8484717 | 100 | 100.148 | 100.148% |  |
| 150% | 10160609 | 150 | 150.091 | 100.060% |  |

- The percentage recovery was found to be within the limit (97-103%).

The results obtained for recovery at 50%, 100%, 150% are within the limits. Hence method is accurate.

Robustness

Table 10: System suitability results for Dabrafenib:

| S.No | Flow Rate (ml/min) | System Suitability Results |  |
| --- | --- | --- | --- |
|  |  | USP Plate Count | USP Tailing |
| 1 | 0.9 | 5784.6 | 1.06 |
| 2 | 1.0 | 5685.4 | 1.08 |
| 3 | 1.1 | 5869.5 | 1.09 |

* Results for actual flow (1.0 ml/min) have been considered from Assay standard.

Table 11: System suitability results for Trametinib:

| S.No | Flow Rate (ml/min) | System Suitability Results |  |
| --- | --- | --- | --- |
|  |  | USP Plate Count | USP Tailing |
| 1 | 0.9 | 6698.3 | 1.46 |
| 2 | 1.0 | 6895.7 | 1.42 |
| 3 | 1.1 | 6983.6 | 1.49 |

* Results for actual flow (1.0ml/min) have been considered from Assay standard.

CONCLUSION

The study is focused to develop and validate HPLC methods for estimation of Dabrafenib and Trametinib in bulk and tablet dosage form. For routine analytical purpose it is desirable to establish methods capable of analyzing huge number of samples in a short time period with good robustness, accuracy and precision without any prior separation steps. HPLC method generates large amount of quality data, which serve as highly powerful and convenient analytical tool. The method shows good reproducibility and good recovery. From the specificity studies, it was found that the developed methods were specific for Dabrafenib and Trametinib

ACKNOWLEDGEMENT

Thе Authors arе thankful to the Management and Principal, Department of Pharmacy, Samskruti college of pharmacy in Ghatkesar, Telangana,for extending support to carry out the research work. Finally, the authors express their gratitude to the Sura Labs, Dilsukhnagar, Hyderabad, for providing research equipment and facilities.

REFERENCES

1.	Chatwal G, Anand SK. Instrumental methods of chemical analysis. 5th ed. New Delhi: Himalaya publishing house; 2002. p. 1.1-8, 2.566-70.

2.	Skoog DA, Holler J, Nieman TA. Principle of instrumental analysis. 5th ed, Saunders college publishing; 1998. p. 778-87.

3.	Skoog, Holler, Nieman. Principals of instrumental analysis. 5th ed, Harcourt Publishers international company; 2001. p. 543-54.

4.	Kemp W. Organic spectroscopy. New York: Palgrave; 2005. p. 7-10, 328-30.

5.	Sethi PD. HPLC: quantitative analysis pharmaceutical formulations, CBS publishers and distributors. New Delhi, India; 2001. p. 3-137.

6.	Michael E, Schartz IS, Krull. Analytical method development and validation; 2004. p. 25-46.

7.	Sharma BK. Instrumental methods of chemical analysis, Introduction to analytical chemistry. 23rd ed.Goel publishing house Meerut; 2004. p. 12-23.

8.	Willard HH, Merritt LL, Dean JA, Settle FA. Instrumental methods of analysis. 7th ed, CBS publishers and distributors. New Delhi; 1986. p. 518-21, 580-610.

9.	Adamovies J. Chromatographic analysis of pharmaceutical. 2nd ed. New York: Marcel Dekker Inc. p. 74, 5-15.

10.	Snyder R, Kirkland J, Glajch L. Practical HPLC method development. 2nd ed. A Wiley international publication; 1997. p. 235, 266-8, 351-353.653-600.686-695.

11.	Basic education in analytical chemistry. Anal Sci. 2001;17(1).

12.	Method validation guidelines International Conference on Harmonization; GENEVA; 1996.

13.	Berry RI, Nash AR. Pharmaceutical process validation, Analytical method validation, Marcel Dekker Inc. New Work. 1993;57:411-28.

14.	Moffat AC, Osselton MD, Widdop B. Clarke’s analysis of drugs and poisons. Vol. 2004. London: Pharmaceutical press; 1601-1602. p. 1109-10.

15.	Florey K. Analysis profile of drugs substances. New York: Academic press; 2005. p. 406-35.

16.	Arora PN, Malhan PK. Biostatistics, Himalaya Publishers house. India. p. 113, 139-40, 154.

17.	Doserge, Wilson and Gisvold’s textbook of organic medicinal and pharmaceutical chemistry. 8th ed. Lippincott Company; 1982. p. 183-97.

18.	Available from: https://www.drugbank.ca/drugs/DB08912.

19.	Available from: https://pubchem.ncbi.nlm.nih.gov/compound/Dabrafenib.

20.	Available from: https://en.wikipedia.org/wiki/Dabrafenib.  Wikipedia.